ACD CDR RFA #4

Requestor:	Joseph Bolek	301-286-1390		GSFC/424

Subject Area:	Contamination

Specific Request:

Delineate plans and procedures for preventing helium contamination of the photomultiplier tubes (PMT’s) during ACD, LAT, and spacecraft integration.  These should include:

- Sampling of purge gases to ensure no helium

- Use of special purge lines (helium can pass through Tygon tubing).

- Actions to be taken if the helium detector alarm goes off.

- Plans for replacement detectors to be in place if the primary detector has to be removed.

Rationale:

The PMT’s are sensitive to helium contamination.  The sensitivity was noted to be at the 5 ppm level, which is near ambient conditions.  Protecting the PMT’s from contamination will be a challenge especially at the instrument and spacecraft level.