## Time history
Run Condition | Time | Temperature (°C) | n disc. Total | n mask. Total | n disc. Y12 | new disc. Y12 | n disc. X7 | new disc. X7 | Tot - Y12,X7
Pisa test |  | 21 | 99 |  | 28 | 0 | 9 | 0 | 62
Post VibeTest |  | 21 | 98 | 39 | 28 | 0 | 9 | 0 | 61
Vacuum, room temp. |  | 21 | 101 | 50 | 28 | 0 | 12 | 3 | 61
HOT Turn On |  | 45 | 97 | 23 | 29 | 2 | 7 | 0 | 61
1st COLD Plateau |  | -15 | 117 | 20 | 29 | 1 | 26 | 16 | 62
1st HOT Plateau |  | 45 | 103 | 27 | 30 | 1 | 10 | 0 | 63
2nd COLD Plateau |  | -15 | 130 | 21 | 37 | 6 | 30 | 2 | 63
2nd HOT Plateau |  | 45 | 111 | 17 | 33 | 0 | 13 | 0 | 65
3rd COLD Plateau |  | -15 | 136 | 21 | 38 | 1 | 34 | 3 | 64
3dr HOT Plateau |  | 45 | 114 | 19 | 36 | 0 | 13 | 0 | 65
4th COLD Plateau |  | -15 | 148 | 26 | 40 | 2 | 38 | 4 | 70
4th HOT Plateau |  | 45 | 122 | 16 | 40 | 0 | 13 | 0 | 69
Vacuum, room temp. |  | 21 | 140 | 16 | 41 | 1 | 28 | 0 | 71
last update: Carmelo Sgro' - 6 Feb 2005

## Disconnected chans Data
[5 empty] | Pisa CPT (before shipment to Alenia) |  |  |  | Post Vibe-Test CPT |  |  |  | CPT in vacuum @ room temperature |  |  |  | CPT @ HOT Turn On |  |  |  | CPT @ 1st COLD plateau |  |  |  | CPT @ 1st HOT plateau |  |  |  | CPT @ 2nd COLD plateau |  |  |  | CPT @ 2nd HOT plateau |  |  |  | CPT @ 3rd COLD plateau |  |  |  | CPT @ 3rd HOT plateau |  |  |  | CPT @ 4th COLD plateau |  |  |  | CPT @ 4th HOT plateau (no data taking) |  |  |  | CPT @ room temperature (in vac.)
(GTCCs, GTRC) | Layer | Tray | Side | MCM | n disc. ch. | n bad ch. | n masked ch. | n broken ch. | n disc. ch. | n bad ch. | n masked ch. | n broken ch. | n disc. ch. | n bad ch. | n masked ch. | n broken ch. | n disc. ch. | n bad ch. | n masked ch. | n broken ch. | n disc. ch. | n bad ch. | n masked ch. | n broken ch. | n disc. ch. | n bad ch. | n masked ch. | n broken ch. | n disc. ch. | n bad ch. | n masked ch. | n broken ch. | n disc. ch. | n bad ch. | n masked ch. | n broken ch. | n disc. ch. | n bad ch. | n masked ch. | n broken ch. | n disc. ch. | n bad ch. | n masked ch. | n broken ch. | n disc. ch. | n bad ch. | n masked ch. | n broken ch. | n disc. ch. | n bad ch. | n masked ch. | n broken ch. | n disc. ch. | n bad ch. | n masked ch. | n broken ch.
[5 empty] | (Low ENC) | (No Counts) |  |  | (Low ENC) | (No Counts) |  |  | (Low ENC) | (No Counts) |  |  | (Low ENC) | (No Counts) |  |  | (Low ENC) | (No Counts) |  |  | (Low ENC) | (No Counts) |  |  | (Low ENC) | (No Counts) |  |  | (Low ENC) | (No Counts) |  |  | (Low ENC) | (No Counts) |  |  | (Low ENC) | (No Counts) |  |  | (Low ENC) | (No Counts) |  |  | (Low ENC) | (No Counts) |  |  | (Low ENC) | (No Counts)
(4, 5, 8) | Y17 | TOP 004 | Back | 11451 | 4 | - | - | - | 4 | 4 | 2 | - | 4 | 4 | 2 | - | 4 | 4 | 0 | - | 4 | 4 | 0 | - | 4 | 4 | 0 | - | 4 | 4 | 0 | - | 4 | 4 | 0 | - | 4 | 4 | 0 | - | 4 | 4 | 0 | - | 4 | 4 | 0 | - | 4 | - | 0 | - | 4 | 4 | 0 | -
(6, 7, 8) | X17 | MID 065 | Front | 11042 | 1 | - | - | - | 1 | 1 | 4 | - | 1 | 1 | 2 | - | 1 | 1 | 1 | - | 0 | 1 | 0 | - | 1 | 1 | 1 | - | 1 | 1 | 0 | - | 1 | 1 | 0 | - | 1 | 1 | 1 | - | 1 | 1 | 0 | - | 1 | 1 | 0 | - | 1 | - | 0 | - | 1 | 1 | 0 | -
(2, 3, 8) | X16 |  | Back | 606 | 1 | - | - | - | 1 | 1 | 0 | - | 1 | 1 | 2 | - | 1 | 1 | 0 | - | 1 | 1 | 0 | - | 1 | 1 | 0 | - | 1 | 1 | 0 | - | 1 | 1 | 0 | - | 1 | 1 | 0 | - | 1 | 1 | 0 | - | 1 | 1 | 0 | - | 1 | - | 0 | - | 1 | 1 | 0 | -
(0, 1, 8) | Y16 | MID 078 | Front | 475 | 0 | - | - | - | 0 | 0 | 3 | - | 0 | 0 | 2 | - | 0 | 0 | 1 | - | 0 | 0 | 0 | - | 0 | 0 | 1 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 2 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | - | 0 | - | 0 | 0 | 0 | -
(4, 5, 7) | Y15 |  | Back | 11523 | 2 | - | - | - | 2 | 3 | 6 | - | 2 | 3 | 0 | - | 2 | 3 | 0 | - | 2 | 3 | 0 | - | 2 | 3 | 0 | - | 2 | 3 | 0 | - | 2 | 3 | 0 | - | 2 | 3 | 0 | - | 2 | 3 | 0 | - | 2 | 3 | 0 | - | 2 | - | 0 | - | 2 | 3 | 0 | -
(6, 7, 7) | X15 | MID 035 | Front | 11480 | 0 | - | - | - | 0 | 0 | 1 | - | 0 | 0 | 0 | - | 0 | 0 | 1 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | - | 0 | - | 0 | 0 | 0 | -
(2, 3, 7) | X14 |  | Back | 11292 | 0 | - | - | - | 0 | 0 | 1 | - | 0 | 0 | 2 | - | 0 | 0 | 1 | - | 0 | 0 | 0 | - | 0 | 0 | 1 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | - | 0 | - | 0 | 0 | 0 | -
(0, 1, 7) | Y14 | MID 071 | Front | 11418 | 2 | - | - | - | 2 | 6 | 2 | - | 2 | 5 | 1 | - | 2 | 5 | 0 | - | 2 | 6 | 0 | - | 2 | 5 | 1 | - | 2 | 6 | 0 | - | 2 | 5 | 0 | - | 2 | 6 | 0 | - | 2 | 5 | 0 | - | 3 | 6 | 0 | - | 2 | - | 0 | - | 2 | 5 | 0 | -
(4, 5, 6) | Y13 |  | Back | 389 | 1 | - | - | - | 1 | 0 | 1 | - | 1 | 0 | 3 | - | 1 | 0 | 3 | - | 1 | 0 | 1 | - | 0 | 0 | 1 | - | 0 | 0 | 1 | - | 0 | 0 | 1 | - | 0 | 0 | 1 | - | 0 | 0 | 1 | - | 0 | 0 | 2 | - | 0 | - | 1 | - | 0 | 0 | 1 | -
(6, 7, 6) | X13 | MID 032 | Front | 11501 | 0 | - | - | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | - | 0 | - | 0 | 0 | 0 | -
(2, 3, 6) | X12 |  | Back | 11422 | 0 | - | - | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 1 | - | 0 | 0 | 0 | - | 0 | 0 | 1 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | - | 1 | - | 0 | 0 | 1 | -
(0, 1, 6) | Y12 | MID 043 | Front | 11457 | 28 | - | - | - | 28 | 28 | 8 | - | 28 | 28 | 15 | - | 29 | 29 | 0 | - | 29 | 31 | 0 | - | 30 | 30 | 0 | - | 37 | 37 | 0 | - | 33 | 33 | 0 | - | 38 | 38 | 0 | - | 36 | 36 | 0 | - | 40 | 40 | 1 | - | 40 | - | 0 | - | 41 | 41 | 0 | -
(4, 5, 5) | Y11 |  | Back | 383 | 0 | - | - | - | 0 | 0 | 0 | - | 0 | 0 | 1 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 1 | - | 0 | - | 0 | - | 0 | 0 | 0 | -
(6, 7, 5) | X11 | MID 025 | Front | 11391 | 3 | - | - | - | 3 | 5 | 3 | - | 3 | 4 | 3 | - | 3 | 4 | 3 | - | 3 | 4 | 4 | - | 3 | 4 | 4 | - | 3 | 4 | 4 | - | 4 | 4 | 4 | - | 3 | 4 | 3 | - | 3 | 4 | 4 | - | 3 | 4 | 3 | - | 4 | - | 3 | - | 3 | 5 | 3 | -
(2, 3, 5) | X10 |  | Back | 11321 | 7 | - | - | - | 7 | 7 | 1 | - | 7 | 7 | 1 | - | 7 | 7 | 1 | - | 7 | 7 | 1 | - | 7 | 7 | 1 | - | 7 | 7 | 1 | - | 7 | 7 | 1 | - | 7 | 7 | 2 | - | 7 | 7 | 1 | - | 7 | 7 | 1 | - | 7 | - | 1 | - | 7 | 7 | 1 | -
(0, 1, 5) | Y10 | MID 051 | Front | 715 | 0 | - | - | - | 0 | 0 | 1 | - | 0 | 0 | 1 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 2 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 1 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | - | 0 | - | 0 | 0 | 0 | -
(4, 5, 4) | Y9 |  | Back | 11392 | 16 | - | - | - | 16 | 16 | 1 | - | 16 | 16 | 4 | - | 15 | 15 | 0 | - | 16 | 16 | 0 | - | 17 | 17 | 0 | - | 17 | 17 | 0 | - | 17 | 17 | 0 | - | 18 | 18 | 0 | - | 17 | 17 | 0 | - | 19 | 19 | 0 | - | 18 | - | 0 | - | 19 | 19 | 0 | -
(6, 7, 4) | X9 | MID 044 | Front | 803 | 2 | - | - | - | 2 | 2 | 1 | - | 2 | 2 | 1 | - | 2 | 2 | 1 | - | 2 | 2 | 0 | - | 2 | 2 | 1 | - | 2 | 2 | 1 | - | 2 | 0 | 1 | - | 2 | 2 | 2 | - | 2 | 2 | 1 | - | 2 | 2 | 2 | - | 2 | - | 2 | - | 2 | 2 | 2 | -
(2, 3, 4) | X8 |  | Back | 11511 | 2 | - | - | - | 2 | 2 | 2 | - | 2 | 2 | 0 | - | 2 | 2 | 0 | - | 2 | 2 | 0 | - | 2 | 2 | 0 | - | 0 | 2 | 0 | - | 2 | 2 | 0 | - | 0 | 2 | 0 | - | 2 | 2 | 0 | - | 2 | 2 | 0 | - | 2 | - | 0 | - | 2 | 2 | 0 | -
(0, 1, 4) | Y8 | MID 031 | Front | 595 | 0 | - | - | - | 0 | 1 | 0 | - | 0 | 1 | 0 | - | 0 | 1 | 0 | - | 0 | 1 | 0 | - | 0 | 1 | 0 | - | 0 | 1 | 0 | - | 0 | 1 | 0 | - | 0 | 1 | 0 | - | 0 | 1 | 0 | - | 0 | 1 | 0 | - | 0 | - | 0 | - | 0 | 1 | 0 | -
(4, 5, 3) | Y7 |  | Back | 853 | 1 | - | - | - | 1 | 1 | 1 | - | 1 | 1 | 1 | - | 1 | 1 | 2 | - | 1 | 1 | 0 | - | 1 | 1 | 1 | - | 1 | 1 | 0 | - | 1 | 1 | 0 | - | 1 | 1 | 3 | - | 1 | 1 | 0 | - | 1 | 1 | 0 | - | 1 | - | 1 | - | 1 | 1 | 1 | -
(6, 7, 3) | X7 | MID 023 | Front | 11022 | 9 | - | - | - | 9 | 9 | 1 | - | 12 | 12 | 3 | - | 7 | 7 | 0 | - | 26 | 25 | 0 | - | 10 | 10 | 0 | - | 30 | 30 | 1 | - | 13 | 13 | 1 | - | 34 | 36 | 0 | - | 13 | 13 | 1 | - | 38 | 36 | 1 | - | 13 | - | 0 | - | 28 | 28 | 0 | -
(2, 3, 3) | X6 |  | Back | 11417 | 3 | - | - | - | 3 | 3 | 0 | - | 3 | 3 | 1 | - | 3 | 3 | 0 | - | 3 | 3 | 0 | - | 3 | 3 | 1 | - | 3 | 3 | 0 | - | 3 | 3 | 0 | - | 3 | 3 | 0 | - | 3 | 3 | 2 | - | 3 | 3 | 0 | - | 3 | - | 0 | - | 3 | 3 | 0 | -
(0, 1, 3) | Y6 | HVY 036 | Front | 740 | 0 | - | - | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | - | 0 | - | 0 | 0 | 0 | -
(4, 5, 2) | Y5 |  | Back | 1087 | 4 | - | - | - | 4 | 2 | 0 | - | 4 | 4 | 0 | - | 4 | 4 | 0 | - | 4 | 4 | 0 | - | 4 | 4 | 1 | - | 4 | 4 | 0 | - | 4 | 4 | 0 | - | 4 | 4 | 0 | - | 4 | 4 | 0 | - | 4 | 4 | 0 | - | 4 | - | 0 | - | 4 | 4 | 0 | -
(6, 7, 2) | X5 | HVY 037 | Front | 396 | 2 | - | - | ~60 | 1 | 2 | 4 | 60 | 1 | 1 | 1 | ~55 | 1 | 2 | 2 | ~61 | 1 | 1 | 4 | ~42 | 1 | 2 | 3 | ~59 | 1 | 1 | 5 | ~50 | 1 | 2 | 1 | ~61 | 1 | 1 | 2 | ~53 | 1 | 2 | 1 | ~63 | 2 | 1 | 3 | ~60 | 2 | - | 1 | - | 1 | 1 | 1 | ~59
(2, 3, 2) | X4 |  | Back | 237 | 0 | - | - | - | 0 | 0 | 3 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | - | 0 | - | 0 | 0 | 0 | -
(0, 1, 2) | Y4 | HVY 018 | Front | 11383 | 0 | - | - | - | 0 | 0 | 0 | - | 0 | 0 | 1 | - | 0 | 0 | 0 | - | 0 | - | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 1 | - | 0 | - | 0 | - | 0 | 0 | 0 | -
(4, 5, 1) | Y3 |  | Back | 1012 | 8 | - | - | - | 8 | 8 | 6 | - | 8 | 8 | 8 | - | 8 | 8 | 6 | - | 10 | 10 | 7 | - | 9 | 9 | 7 | - | 12 | 12 | 6 | - | 11 | 11 | 6 | - | 12 | 12 | 6 | - | 12 | 12 | 6 | - | 13 | 13 | 6 | - | 12 | - | 5 | - | 15 | 13 | 5 | -
(6, 7, 1) | X3 | HVY 005 | Front | 11157 | 1 | - | - | - | 1 | 1 | 0 | - | 1 | 1 | 0 | - | 1 | 1 | 0 | - | 1 | 1 | 0 | - | 1 | 1 | 0 | - | 1 | 1 | 0 | - | 1 | 1 | 0 | - | 1 | 1 | 0 | - | 1 | 1 | 0 | - | 1 | 1 | 1 | - | 1 | - | 0 | - | 1 | 1 | 0 | -
(2, 3, 1) | X2 |  | Back | 1118 | 1 | - | - | - | 1 | 2 | 2 | - | 1 | 1 | 3 | - | 2 | 2 | 3 | - | 0 | 2 | 2 | - | 2 | 2 | 2 | - | 1 | 2 | 1 | - | 1 | 2 | 1 | - | 1 | 1 | 0 | - | 1 | 2 | 1 | - | 1 | 1 | 0 | - | 2 | - | 0 | - | 2 | 2 | 0 | -
(0, 1, 1) | Y2 | LGT 017 | Front | 966 | 0 | - | - | - | 0 | 0 | 0 | - | 0 | 0 | 1 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | - | 0 | - | 0 | 0 | 0 | -
(4, 5, 0) | Y1 |  | Back | 965 | 1 | - | - | - | 1 | 1 | 0 | - | 1 | 1 | 0 | - | 1 | 1 | 0 | - | 1 | 1 | 0 | - | 1 | 1 | 0 | - | 1 | 1 | 0 | - | 1 | 1 | 0 | - | 1 | 1 | 0 | - | 1 | 1 | 0 | - | 1 | 1 | 1 | - | 1 | - | 0 | - | 1 | 1 | 0 | -
(6, 7, 0) | X1 | LGT 004 | Front | 662 | 8 | - | - | - | 8 | 8 | 1 | - | 8 | 8 | 1 | - | 8 | 8 | 1 | - | 8 | 8 | 1 | - | 8 | 8 | 1 | - | 8 | 8 | 1 | - | 8 | 8 | 1 | - | 8 | 8 | 1 | - | 8 | 8 | 1 | - | 8 | 8 | 1 | - | 8 | - | 1 | - | 8 | 8 | 1 | -
(2, 3, 0) | X0 |  | Back | 480 | 0 | - | - | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | - | 0 | - | 0 | 0 | 0 | -
(0, 1, 0) | Y0 | BTM 006 | Front | 11140 | 0 | - | - | - | 0 | 0 | 1 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 0 | - | 0 | 0 | 2 | - | 0 | - | 0 | - | 0 | 0 | 0 | -
|  | Total |  |  | 99 |  |  |  | 98 | 104 | 39 |  | 101 | 105 | 50 |  | 97 | 102 | 23 |  | 117 | 125 | 20 |  | 103 | 109 | 27 |  | 130 | 139 | 21 |  | 111 | 115 | 17 |  | 136 | 146 | 21 |  | 114 | 121 | 19 |  | 148 | 150 | 26 |  | 122 | 0 | 16 |  | 140 | 144 | 16
Note:
| Used Test |  | Description
n disc. ch. (Low ENC) | NoiseAndGain |  | ENC < 500 e-
n bad ch. (No Counts) | DataTaking |  | Exactly 0 counts
n broken ch. | DataTaking |  | Strip broken along the ladder: half count with respect to full strip

## MID 043 Front
Post Vibe-Test CPT (308000837) | CPT in vacuum @ room temperature (308000862) | CPT @ HOT Turn On (308000872) | CPT @ 1st COLD plateau (308000882) | CPT @ 1st HOT plateau (308000908) | CPT @ 2nd COLD plateau (308000925) | CPT @ 2nd HOT plateau (308000944) | CPT @ 3rd COLD plateau (308000963) | CPT @ 3rd HOT plateau (308000983) | CPT @ 4th COLD plateau (3080001000) | CPT @ 4th HOT plateau (3080001015) | CPT @ room temperature in vac. (3080001024)
28 disconnected channels: | 28 disconnected channels: | 29 disconnected channels: | 29 disconnected channels: | 30 disconnected channels: | 37 disconnected channels: | 33 disconnected channels: | 38 disconnected channels: | 36 disconnected channels: | 40 disconnected channels: | 40 disconnected channels: | 41 disconnected channels:
[821, 832, 961, 964, 969, 995, 996, 1107, 1110, 1123, 1133, 1135, 1136, 1138, 1140, 1142, 1144, 1146, 1148, 1152, 1159, 1175, 1177, 1195, 1216, 1217, 1227, 1268] | [821, 832, 961, 964, 969, 995, 996, 1107, 1110, 1123, 1133, 1135, 1136, 1138, 1140, 1142, 1144, 1146, 1148, 1152, 1159, 1175, 1177, 1195, 1216, 1217, 1227, 1268] | [821, 832, 846, 961, 964, 969, 995, 996, 1107, 1110, 1123, 1129, 1133, 1136, 1138, 1140, 1142, 1144, 1146, 1148, 1152, 1159, 1175, 1177, 1195, 1216, 1217, 1227, 1268] | [782, 821, 832, 961, 964, 969, 995, 996, 1107, 1110, 1123, 1133, 1136, 1138, 1140, 1142, 1144, 1146, 1148, 1152, 1159, 1164, 1175, 1177, 1195, 1216, 1217, 1227, 1268] | [821, 832, 846, 961, 964, 969, 995, 996, 1107, 1110, 1123, 1129, 1133, 1135, 1136, 1138, 1140, 1142, 1144, 1146, 1148, 1152, 1159, 1175, 1177, 1195, 1216, 1217, 1227, 1268] | [782, 821, 832, 846, 923, 961, 964, 969, 994, 995, 996, 1107, 1110, 1115, 1123, 1129, 1133, 1135, 1136, 1138, 1140, 1142, 1144, 1145, 1146, 1148, 1150, 1152, 1159, 1164, 1175, 1177, 1195, 1216, 1217, 1227, 1268] | [782, 821, 832, 846, 961, 964, 969, 994, 995, 996, 1107, 1110, 1123, 1129, 1133, 1135, 1136, 1138, 1140, 1142, 1144, 1146, 1148, 1152, 1159, 1164, 1175, 1177, 1195, 1216, 1217, 1227, 1268] | [782, 821, 832, 839, 846, 923, 961, 964, 969, 994, 995, 996, 1107, 1110, 1115, 1123, 1129, 1133, 1135, 1136, 1138, 1140, 1142, 1144, 1145, 1146, 1148, 1150, 1152, 1159, 1164, 1175, 1177, 1195, 1216, 1217, 1227, 1268] | [782, 821, 832, 846, 923, 961, 964, 969, 994, 995, 996, 1107, 1110, 1115, 1123, 1129, 1133, 1135, 1136, 1138, 1140, 1142, 1144, 1146, 1148, 1150, 1152, 1159, 1164, 1175, 1177, 1195, 1216, 1217, 1227, 1268] | [782, 821, 829, 832, 839, 846, 923, 961, 964, 969, 994, 995, 996, 1103, 1107, 1110, 1115, 1123, 1129, 1133, 1135, 1136, 1138, 1140, 1142, 1144, 1145, 1146, 1148, 1150, 1152, 1159, 1164, 1175, 1177, 1195, 1216, 1217, 1227, 1268] | [782, 821, 829, 832, 839, 846, 923, 961, 964, 969, 994, 995, 996, 1103, 1107, 1110, 1115, 1123, 1129, 1133, 1135, 1136, 1138, 1140, 1142, 1144, 1145, 1146, 1148, 1150, 1152, 1159, 1164, 1175, 1177, 1195, 1216, 1217, 1227, 1268] | [782, 821, 829, 832, 839, 846, 923, 961, 964, 969, 994, 995, 996, 1103, 1107, 1110, 1115, 1123, 1129, 1133, 1135, 1136, 1137, 1138, 1140, 1142, 1144, 1145, 1146, 1148, 1150, 1152, 1159, 1164, 1175, 1177, 1195, 1216, 1217, 1227, 1268]
Note:
In red those channels that appear disconnected for the first time

## MID 023 Front
Post Vibe-Test CPT (308000837) | CPT in vacuum @ room temperature (308000862) | CPT @ HOT Turn On (308000872) | CPT @ 1st COLD plateau (308000882) | CPT @ 1st HOT plateau (308000908) | CPT @ 2nd COLD plateau (308000925) | CPT @ 2nd HOT plateau (308000944) | CPT @ 3rd COLD plateau (308000963) | CPT @ 3rd HOT plateau (308000983) | CPT @ 4th COLD plateau (3080001000) | CPT @ 4th HOT plateau (3080001015) | CPT @ room temperature in vac. (3080001024)
9 disconnected channels: | 12 disconnected channels | 7 disconnected channels | 26 disconnected channels | 10 disconnected channels | 30 disconnected channels | 13 disconnected channels | 34 disconnected channels | 13 disconnected channels: | 38 disconnected channels: | 13 disconnected channels: | 28 disconnected channels:
[10, 123, 1164, 1284, 1285, 1286, 1306, 1475, 1492] | [10, 123, 229, 232, 238, 1164, 1284, 1285, 1286, 1306, 1475, 1492] | [10, 123, 1285, 1286, 1306, 1475, 1492] | [10, 30, 39, 40, 41, 42, 123, 215, 220, 223, 226, 228, 232, 237, 238, 241, 244, 248, 1164, 1190, 1284, 1285, 1286, 1306, 1475, 1489] | [10, 123, 1190, 1284, 1285, 1286, 1306, 1475, 1489, 1492] | [10, 30, 39, 40, 41, 42, 123, 215, 220, 223, 225, 226, 228, 229, 232, 237, 238, 241, 244, 248, 1164, 1190, 1284, 1285, 1286, 1306, 1398, 1475, 1489, 1492] | [10, 123, 232, 1164, 1190, 1284, 1285, 1286, 1306, 1398, 1475, 1489, 1492] | [10, 30, 39, 40, 41, 42, 123, 150, 215, 220, 223, 225, 226, 228, 229, 232, 236, 237, 238, 241, 243, 244, 248, 1163, 1164, 1190, 1284, 1285, 1286, 1306, 1398, 1475, 1489, 1492] | [10, 123, 232, 1164, 1190, 1284, 1285, 1286, 1306, 1398, 1475, 1489, 1492] | [10, 30, 39, 40, 41, 42, 43, 49, 123, 150, 212, 215, 220, 223, 225, 226, 228, 229, 232, 236, 237, 238, 239, 241, 243, 244, 248, 1163, 1164, 1190, 1284, 1285, 1286, 1306, 1398, 1475, 1489, 1492] | [10, 123, 1163, 1164, 1190, 1284, 1285, 1286, 1306, 1398, 1475, 1489, 1492] | [10, 30, 41, 42, 123, 212, 215, 220, 223, 226, 228, 229, 232, 236, 238, 243, 248, 1163, 1164, 1190, 1284, 1285, 1286,1306, 1398, 1475, 1489, 1492]
Note:
In red those channels that appear disconnected for the first time